## WSS-09
Rain Gauge
Default RS485 Address: 0x01
Baud Rate: | For WSS-09:4800
Rainfall query:
master→slave
address | function code | start register address high | start register address low | Register length high | Register length is low | CRC16 Low | CRC16 High
0X01 | 0X03 | 0X01 | 0XF4 | 0X00 | 0X0C | 0XC1 | 0X05
slave → master
Address |  |  | 0X01
Function code |  |  | 0X03
Data length |  |  | 0X18
Register 0 data high |  |  | 0X00 | wind speed：3.6m/s
Register 0 data low |  |  | 0X7D
Register 1 data high |  |  | 0X00 | wind scale：Level 3
Register 1 data low |  |  | 0X03
Register 2 data high |  |  | 0X00 | Wind direction : 1 (Northeast)
Register 2 data low |  |  | 0X01
Register 3 data high |  |  | 0X01 | wind angle：45.0°
Register 3 data low |  |  | 0XC2
Register 4 data high |  |  | 0X02 | Humidity: 65.8%RH
Register 4 data low |  |  | 0X92
Register 5 data high |  |  | 0XFF | Temperature: -10.1℃
Register 5 data low |  |  | 0X9B
Register 6 data high |  |  | 0X02 | Noise: 51.0dB
Register 6 data low |  |  | 0X26
Register 7 data high |  |  | 0X00 | PM2.5: 90μg/m3
Register 7 data low |  |  | 0X62
Register 8 data high |  |  | 0X01 | PM10: 288μg/m3
Register 8 data low |  |  | 0X20
Register 9 data high |  |  | 0X03 | Atmospheric pressure: 101.3kPa‌
Register 9 data low |  |  | 0XF5
Register 10 data high |  |  | 0X00 | Light Illuminance: 61235 Lux
Register 10 data low |  |  | 0X00
Register 11 data high |  |  | 0XEF
Register 11 data low |  |  | 0X33
CRC16 low |  |  | 0XBD
CRC16 high |  |  | 0X13

## WSS-21
Rain Gauge
Default RS485 Address: 0x06
Baud Rate: | For WSS-21:4800
For WSS-01:9600
Modification Address:
master→slave
Original address | Function code | Register address high | Register address low | Start address high | Start address low | CRC16 low | CRC16 high
0X06 | 0X06 | 0X00 | 0X30 | 0X00 | 0X02 | 0X09 | 0XB3
slave → master
Return with the same path
Query / set time:
master→slave
Original address | Function code | Register address high | Register address low | Register length high | Register length low | CRC16 low | CRC16 high
0X06 | 0X03 | 0X00 | 0X34 | 0X00 | 0X03 | 0X45 | 0XB2
slave → master
Address | Function code | Data length | Year | Month | Day | Hour | Minutes | Second | CRC16 low | CRC16 high
0X06 | 0X03 | 0X06 | 0X20 | 0X03 | 0X30 | 0X10 | 0X25 | 0X10 | 51 | BC
BCD code, indicating: 10:25:10 on March 30, 20
If the clock has deviation, the clock can be calibrated, and the host → slave
Address |  | 0X06
Function code |  | 0X10
Start register address high |  | 0X00
Start register address low |  | 0X34
Register length high |  | 0X01
Register length low |  | 0X03
Data length |  | 0X06
Year |  | 0X20 | BCD code Said: 17:06:28 on April 3, 20
Month |  | 0X04
Day |  | 0X03
Hour |  | 0X17
Minutes |  | 0X06
Second |  | 0X28
CRC16 low |  | 0XE2
CRC16 high |  | 0XF4
Rainfall query:
Original address | Function code | Start register address high | Start register address low | Register length high | Register length low | CRC16 low | CRC16 high
0X06 | 0X03 | 0X00 | 0X00 | 0X00 | 0X0A | 0XC4 | 0X7A
slave → master
Address |  |  | 0X06
Function code |  |  | 0X03
Data length |  |  | 0X14
Register 0 data high |  |  | 0X00 | Daily rainfall: 10.0 mm Rainfall from 0 a.m. to now
Register 0 data low |  |  | 0X64
Register 1 data high |  |  | 0X00 | Instantaneous rainfall: 1.6 mm Rainfall between queries
Register 1 data low |  |  | 0X10
Register 2 data high |  |  | 0X00 | Yesterday's rainfall: 8.0mm Rainfall in 24 hours yesterday
Register 2 data low |  |  | 0X50
Register 3 data high |  |  | 0X06 | Total rainfall: 166.5mm Total rainfall after the sensor is powered on
Register 3 data low |  |  | 0X81
Register 4 data high |  |  | 0X00 | Hourly rainfall: 0.2mm
Register 4 data low |  |  | 0X02
Register 5 data high |  |  | 0X00 | Rainfall in last hour: 0.2mm
Register 5 data low |  |  | 0X02
Register 6 data high |  |  | 0X00 | 24-hour maximum rainfall: 10.0mm
Register 6 data low |  |  | 0X64
Register 7 data high |  |  | 0X01 | 24:00 maximum rainfall period 01:00~02:00
Register 7 data low |  |  | 0X02
Register 8 data high |  |  | 0X00 | 24-hour minimum rainfall: 0.0mm
Register 8 data low |  |  | 0X00
Register 9 data high |  |  | 0X03 | 24-hour minimum rainfall period 03:00~04:00
Register 9 data low |  |  | 0X04
CRC16 low |  |  | 0X24
CRC16 high |  |  | 0XDC
Rainfall data reset setting:
Original address | Function code | Register address high | Register address low | Data content high | Data content low | CRC16 low | CRC16 high
0X06 | 0X06 | 0X00 | 0X37 | 0X00 | 0X03 | 0X79 | 0XB2

## WSS-22
Integrated wind speed and direction sensor
Default RS485 Address: 0x01
Baud Rate: | For WSS-22:4800
For WSS-02:9600
Modification Address:
master→slave
original address | function code | start register high | start register low | start address high | start address low | CRC16 low | CRC16 high
0X07 | 0X06 | 0X00 | 0X30 | 0X00 | 0X02 | 0X08 | 0X62
Query data:
master→slave
address | function code | start register address high | start register address low | Register length high | Register length is low | CRC16 low | CRC16 high
0X07 | 0X03 | 0X00 | 0X00 | 0X00 | 0X04 | 0X44 | 0X6F
slave → master
address |  |  | 0X07
function code |  |  | 0X03
Data length |  |  | 0X08
Register 0 data high |  |  | 0X00 | wind speed：3.6m/s
Register 0 data low |  |  | 0X24
Register 1 data high |  |  | 0X00 | wind scale：Level 3
Register 1 data low |  |  | 0X03
Register 2 data high |  |  | 0X02 | wind angle：66.6°
Register 2 data low |  |  | 0X9A
Register 3 data high |  |  | 0X00 | Wind direction: east-northeast
register 3 data low |  |  | 0X03
CRC16 low |  |  | 0XCA
CRC16 high |  |  | 0XCB
set temperature:
address | function code | start register high | register low address | Register length high | register data low | CRC16 low | CRC16 high
0X07 | 0X06 | 0X00 | 0X6C | 0X01 | 0X2C | 0X49 | 0XFC
set 5℃ temperature:
address | function code | start register high | register low address | Register length high | register data low | CRC16 low | CRC16 high
0X07 | 0X06 | 0X00 | 0X6D | 0X00 | 0X32 | 0X99 | 0XA4
Query start/stop heating temperature data：
master→slave
address | function code | start register address high | start register address low | Register length high | Register length is low | CRC16 low | CRC16 high
0X07 | 0X03 | 0X00 | 0X6C | 0X00 | 0X02 | 0X04 | 0X70
slave → master
address | function code | Data length | Register 0 data high | Register 0 data low | Register 1 data high | Register 1 data low | CRC16 low | CRC16 high
0X07 | 0X03 | 0X04 | 0X01 | 0X2C | 0X00 | 0X32 | 0XDD | 0XD3
|  |  | stop heating temperature：30℃ |  | Start heating temperature：5℃

## WSS-23
CO2+PM2.5+PM10
Default RS485 Address: 0x0C
Baud Rate: | For WSS-23:4800
For WSS-03:9600
Modification Address:
Original address | Function code | Start register address high | Start register address low | Start address high | Start address low | CRC16 low | CRC16 high
0X0C | 0X06 | 0X00 | 0X30 | 0X00 | 0X02 | 0X09 | 0X19
Query data:
master→slave
Address | Function code | Start address high | Start address low | Register length high | Register length low | CRC16 low | CRC16 high
0X0C | 0X03 | 0X00 | 0X30 | 0X00 | 0X03 | 0X04 | 0XD6
slave → master
Address | Function Code | Data Length | Register 0 Data High | Register 0 Data Low | Register 1 Data High | Register 1 Data Low | Register 2 Data High | Register 2 Data Low | CRC16 Low | CRC16 High
0X0C | 0X03 | 0X06 | 0X0E | 0X62 | 0X00 | 0X62 | 0X01 | 0X20 | 0X20 | 0X95
|  |  | CO2 |  | PM2.5 |  | PM10

## WSS-24
Rain and snow sensor
Default RS485 Address: 0x04
Baud Rate: | For WSS-24:4800
For WSS-04:9600
Modification Address:
master→slave
Original address | Function code | Register address high | Register address low | Start address high | Start address low | CRC16 low | CRC16 high
0X04 | 0X06 | 0X00 | 0X30 | 0X00 | 0X02 | 0X08 | 0X51
slave → master
Original address | Function code | Data length | Start address high | Start address low | CRC16 low | CRC16 high
0X04 | 0X06 | 0X02 | 0X00 | 0X02 | 0XF5 | 0X49
Query data:
master→slave
Address | function code | start register address high | start register address low | Register length high | Register length is low | CRC16 low | CRC16 high
0X04 | 0X03 | 0X00 | 0X00 | 0X00 | 0X01 | 0X84 | 0X5F
slave → master
When there is no rain (snow)
Address | Function code | Data length | Register 0 data high |  | Register 0 data low |  | CRC16 low | CRC16 high
0X04 | 0X03 | 0X02 | 0X00 |  | 0X0A |  | 0XF4 | 0X43
No rain (snow)
slave → master
When there is rain (snow)
Address | Function code | Data length | Register 0 data high |  | Register 0 data low |  | CRC16 low | CRC16 high
0X04 | 0X03 | 0X02 | 0X00 |  | 0X0B |  | 0X35 | 0X83
Rain (snow)
Set active reporting:
Address | Function code | Reporting mode high | Reporting mode high | Reporting time high | Reporting time low | CRC16 low | CRC16 high
0X04 | 0X06 | 0X00 | 0X33 | 0X00 | 0X3C | 0X79 | 0X81
|  | 0X0000：Turn off automatic reporting |  | Automatic reporting interval, unit: Second At this time, the upload speed is 60s / time, which is invalid when automatic reporting is turned off
Set the starting temperature of heating:
Address | Function code | start register address high | start register address low | Enable temperature data high | Enable temperature data low | CRC16 low | CRC16 high
0X04 | 0X06 | 0X00 | 0X5A | 0X00 | 0X96 | 0X29 | 0XE2
15.0℃
Set heating stop temperature:
Address | Function code | start register address high | start register address low | Stop temperature data high | Stoptemperature data low | CRC16 low | CRC16 high
0X04 | 0X06 | 0X00 | 0X5B | 0X00 | 0XC8 | 0XF9 | 0XDA
20.0℃

## WSS-25
TEM_HUM_LIGHT_PRESSURE
Default RS485 Address: 0x08
Baud Rate: | For WSS-25:4800
For WSS-05:9600
Modification Address:
master→slave
Original Address | Function Code | Start Register High | Start Register Low | Start Address High | Start Address Low | CRC16 Low | CRC16 High
0X08 | 0X06 | 0X00 | 0X30 | 0X00 | 0X02 | 0X08 | 0X9D
Query Data:
master→slave
Address | Function Code | Start Register Address High | Start Register Address Low | Register Length High | Register Length Low | CRC16 Low | CRC16 High
0X08 | 0X03 | 0X00 | 0X00 | 0X00 | 0X04 | 0X44 | 0X90
slave → master
Address | Function Code | Data Length | Register 0 Data High | Register 0 Data Low | Register 1 Data High | Register 1 Data Low | Register 2 Data High | Register 2 Data Low | Register 3 Data High | Register 3 Data Low | CRC16 Low | CRC16 High
0X0B | 0X03 | 0X08 | 0XFF | 0XDD | 0XD1 | 0X64 | 0X04 | 0XD2 | 0X27 | 0X28 | 0X33 | 0XBC
|  |  | Temperature |  | Humidity |  | Light Illuminance |  | Atmospheric Pressure

## WSS-26
Simple Total Radiation Sensor
Default RS485 Address: 0x0B
Baud Rate: | For WSS-26:4800
For WSS-06:9600
Modification Address:
master→slave
Original Address | Function Code | Register Address High | Register Address Low | Start Address High | Start Address Low | CRC16 Low | CRC16 High
0X0B | 0X06 | 0X00 | 0X30 | 0X00 | 0X02 | 0X08 | 0XAE
Query Data:
master→slave
Address | Function Code | Start Register Address High | Start Register Address Low | Register Length High | Register Length Low | CRC16 Low | CRC16 High
0X0B | 0X03 | 0X00 | 0X00 | 0X00 | 0X01 | 0X84 | 0XA0
slave → master
Address | Function Code | Data Length | Register 0 Data High | Register 0 Data Low | CRC16 Low | CRC16 High
0X0B | 0X03 | 0X02 | 0X30 | 0X39 | 0XF4 | 0X57
Total radiation: 1234.5W/m²

## WSS-27
Photosynthetically Active Radiation Sensor
Default RS485 Address: 0x0A
Baud Rate: | For WSS-27:4800
For WSS-07:9600
Modification Address:
master→slave
Original Address | Function Code | Register Address High | Register Address Low | Start Address High | Start Address Low | CRC16 Low | CRC16 High
0X0A | 0X06 | 0X00 | 0X30 | 0X00 | 0X02 | 0X09 | 0X7F
Query data:
master→slave
address | function code | start register address high | start register address low | Register length high | Register length is low | CRC16 Low | CRC16 High
0X0A | 0X03 | 0X00 | 0X00 | 0X00 | 0X01 | 0X85 | 0X71
slave → master
address | function code | Data length | Register 0 data high | Register 0 data low | CRC16 | CRC16
[5 empty] | low | high
0X0A | 0X03 | 0X02 | 0X00 | 0XC8 | 0X1C | 0X13
Photosynthetically efficient：200μmol/m2•s